## Gold Seal Cat. No.
MR QUICK REFERENCE - 3/3/08
GS Cat. No. | Sketch #
L1000PN | 8-144
L1000PJ | 8-116
L1000EE or L1000EF | 8-136
S1000PN | 8-108

## Moving Metal Lines
MR QUICK REFERENCE - 3/3/08
1.5T Systems
Feet
| X | Y | Z
Objects 100 -400 lbs | 9.12 | 14.66 | 9.12
Cars, Minivans, Vans, Pickup Trucks, Ambulances | 15.5 | 21
Bus, Trucks (Utility, Dump, Semi) | 18.1 | 24.5
Objects > 400 lbs, Elevators, Trains, Subways | Refer to PIM
Meters
Objects 100 -400 lbs | 2.78 | 4.47 | 2.78
Cars, Minivans, Vans, Pickup Trucks, Ambulances | 4.72 | 6.4
Bus, Trucks (Utility, Dump, Semi) | 5.52 | 7.47
Objects > 400 lbs, Elevators, Trains, Subways | Refer to PIM
3.0T
Feet
| X | Y | Z
Objects 100 -400 lbs | 10.1708 | 18.7008 | 10.1708
Cars, Minivans, Vans, Pickup Trucks, Ambulances | 21 | 26
Bus, Trucks (Utility, Dump, Semi) | 24.5 | 30.3
Objects > 400 lbs, Elevators, Trains, Subways | Refer to PIM
Meters
Objects 100 -400 lbs | 3.1 | 5.7 | 3.1
Cars, Minivans, Vans, Pickup Trucks, Ambulances | 6.4 | 7.92
Bus, Trucks (Utility, Dump, Semi) | 7.47 | 9.25
Objects > 400 lbs, Elevators, Trains, Subways | Refer to PIM
Openspeed
Radius in Feet
| HWY | PRP
Objects 100 -400 lbs | Refer to PIM
Cars, Minivans, Vans, Pickup Trucks, Ambulances | 31.5 | 24.9
Bus, Trucks (Utility, Dump, Semi) | 36.7 | 29.2
Objects > 400 lbs, Elevators, Trains, Subways | Refer to PIM
Meters
Objects 100 -400 lbs | Refer to PIM
Cars, Minivans, Vans, Pickup Trucks, Ambulances | 9.6 | 7.6
Bus, Trucks (Utility, Dump, Semi) | 11.2 | 8.9
Objects > 400 lbs, Elevators, Trains, Subways | Refer to PIM
Ovation
Radius in Feet
| HWY | PRP
Objects 100 -400 lbs | Refer to PIM
Cars, Minivans, Vans, Pickup Trucks, Ambulances | 36.0892 | 19.685
Bus, Trucks (Utility, Dump, Semi) | 42.6509 | 22.9659
Objects > 400 lbs, Elevators, Trains, Subways | Refer to PIM
Meters
Objects 100 -400 lbs | Refer to PIM
Cars, Minivans, Vans, Pickup Trucks, Ambulances | 11 | 6
Bus, Trucks (Utility, Dump, Semi) | 13 | 7
Objects > 400 lbs, Elevators, Trains, Subways | Refer to PIM
Profile
Radius in Feet
| No shield |  | 2mm shield
Train | 164.042 |  | 131.2336
Cars, Minivans, Vans, Pickup Trucks, Ambulances | 36.0892 |  | 22.9659
Bus, Trucks (Utility, Dump, Semi) | 42.6509 |  | 26.2467
AC Lines | 16.4042 |  | 13.1234
Objects <181 Kg | 9.8425 |  | Outside magnet room
Elevator | 36.0892 |  | 22.9659
Meters
Train | 50 |  | 40
Cars, Minivans, Vans, Pickup Trucks, Ambulances | 11 |  | 7
Bus, Trucks (Utility, Dump, Semi) | 13 |  | 8
AC Lines | 5 |  | 4
Objects <181 Kg | 3 |  | Outside magnet room
Elevator | 11 |  | 7

## Gauss fields
MR QUICK REFERENCE - 3/3/08
1.5T Systems
Feet
Field | X | Y | Z
0.5 | 12.23 | 21.68 | 12.23
1 | 10.76 | 18.7 | 10.76
3 | 9.12 | 14.66 | 9.12
5 | 8.13 | 13.12 | 8.13
10 | 7.18 | 11.51 | 7.18
30 | 6.2 | 9.25 | 6.2
50 | 5.87 | 8.46 | 5.87
100 | 5.3 | 7.4 | 5.3
200 | 4.84 | 6.51 | 4.84
400 | 4.51 | 5.74 | 4.51
2000 | 3.28 | 4.26 | 3.28
Meters
0.5 | 3.73 | 6.61 | 3.73
1 | 3.28 | 5.7 | 3.28
3 | 2.78 | 4.47 | 2.78
5 | 2.48 | 4 | 2.48
10 | 2.19 | 3.51 | 2.19
30 | 1.89 | 2.82 | 1.89
50 | 1.79 | 2.58 | 1.79
100 | 1.61 | 2.25 | 1.61
200 | 1.47 | 1.98 | 1.47
400 | 1.38 | 1.75 | 1.38
2000 | 1 | 1.3 | 1
3.0T
Feet
Field | X | Y | Z
0.5 | 18.0442 | 29.8558 | 18.0442
1 | 14.4358 | 24.2783 | 14.4358
3 | 10.1708 | 18.7008 | 10.1708
5 | 9.1867 | 16.4042 | 9.1867
10 | 8.53 | 14.4358 | 8.53
30 | 7.5458 | 11.6467 | 7.5458
50 | 7.2175 | 10.4983 | 7.2175
100 | 6.5617 | 9.1867 | 6.5617
200 | 5.9058 | 8.2017 | 5.9058
400 | 5.5775 | 7.2175 | 5.5775
2000 | 4.4292 | 5.4133 | 4.4292
Meters
0.5 | 5.5 | 9.1 | 5.5
1 | 4.4 | 7.4 | 4.4
3 | 3.1 | 5.7 | 3.1
5 | 2.8 | 5 | 2.8
10 | 2.6 | 4.4 | 2.6
30 | 2.3 | 3.55 | 2.3
50 | 2.2 | 3.2 | 2.2
100 | 2 | 2.8 | 2
200 | 1.8 | 2.5 | 1.8
400 | 1.7 | 2.2 | 1.7
2000 | 1.35 | 1.65 | 1.35
Openspeed
Feet
Field | X | Y | Z
0.5 | 30.18 | 30.18 | 34.11
1 | 24.44 | 24.44 | 25.58
3 | 17.88 | 17.88 | 15.25
5 | 15.42 | 15.42 | 12.3
10 | 12.96 | 12.96 | 10.82
30 | 9.84 | 9.84 | 9.35
50 | 8.86 | 8.86 | 8.53
100 | 7.38 | 7.38 | 7.71
200 | 6.4 | 6.4 | 6.72
400 | 5.43 | 5.43 | 5.76
2000 | 3.78 | 3.78 | 3.45
Meters
0.5 | 9.2 | 9.2 | 10.4
1 | 7.45 | 7.45 | 7.8
3 | 5.45 | 5.45 | 4.65
5 | 4.7 | 4.7 | 3.75
10 | 3.95 | 3.95 | 3.3
30 | 3 | 3 | 2.85
50 | 2.7 | 2.7 | 2.6
100 | 2.25 | 2.25 | 2.35
200 | 1.95 | 1.95 | 2.05
400 | 1.65 | 1.65 | 1.75
2000 | 1.15 | 1.15 | 1.05
Ovation
Feet
Field | X | Y | Z
0.5 | 12.3032 | 12.7953 | 15.748
1 | 9.1864 | 8.2021 | 11.1549
3 | 7.71 | 6.5617 | 9.5144
5 | 6.2336 | 5.4134 | 8.2021
10 | 5.5774 | 4.5932 | 6.8898
Meters
0.5 | 3.75 | 3.9 | 4.8
1 | 2.8 | 2.5 | 3.4
3 | 2.35 | 2 | 2.9
5 | 1.9 | 1.65 | 2.5
10 | 1.7 | 1.4 | 2.1
Profile
Feet
Field | X | Y | Z
1 | 9.5144 | 6.2336 | 10.8268
3 | 6.5617 | 5.7415 | 7.874
5 | 5.9055 | 5.2493 | 6.5617
10 | 4.2651 | 4.9213 | 4.9213
Meters
1 | 2.9 | 1.9 | 3.3
3 | 2 | 1.75 | 2.4
5 | 1.8 | 1.6 | 2
10 | 1.3 | 1.5 | 1.5

## Documentation
MR QUICK REFERENCE - 3/3/08
Magnet Manuals | http://www.gehealthcare.com/company/docs/siteplanning.html
MR Coil and Shield Coolers (MRCC) | http://3.87.118.28/optec3/common/excite/content/85417.htm
1.5T Signa Hde | Direction 5143464 rev 04
1.5T Signa HDx | Direction 5159901 rev 04
1.5T Signa Excite HD | Direction 5133301 rev 04
1.5T Signa Twinspeed HDx | Direction 5159902 rev 03
1.5T Signa Twinspeed Excite HD | Direction 5133302 rev 04
3.0T Signa HDx | Direction 5159903 rev 04
3.0T Signa Excite HD | Direction 5133303 rev 05
0.35T Signa Ovation 5 | Direction 5165376 rev 02
0.35T Signa Ovation | Direction 2275665 rev 09
0.7T Signa Openspeed | Direction 2402692 rev 00
0.5T Signa Contour | Direction 2189673 rev 04
0.2T Signa Profile 5 | Direction 5117792 rev 02
0.2T Signa Profile | Direction 2186839 rev 04
0.2T Signa Profile HD | Direction 5196540 rev 01
Signa MR750 3.0T | Direction 5500101 rev 01

## MR750 3T
MR QUICK REFERENCE - 3/3/08
Signa MR750 3.0T | Direction 5500101 rev 01
Magnet Type | Active
Magnet Mounting | Required. RF vendor supplies and installs anchors.
Table Anchored | Mobile Table
Gauss plots | refer to Gauss field tab
Cryogen Vent Size and Location from Iso | 8" Diameter 15.98" left 25.75" back
gantry cable access | Top of Magnet
Magnet weight | 24,952 lbs
conduit sizes -OC to Equip | (1) 2" & (1) 3"
RF Screen Room | 100db (10-150 MHz planewave)
Cable Kits
M7000BN | L1 = Short L2 = Short L3, L4, L5 & L6 = One length
M7000BP | L1 = Short L2 = Long L3, L4, L5 & L6 = One length
M7000BR | L1 = Long L2 = Short L3, L4, L5 & L6 = One length
voltage/phase MDP cat NO. | 480V / 3Phase
recommended ceiling height | 8'-9"
absolute min. ceiling height | 8'-2 1/2" With low ceiling height kit
absolute min. Magnet room size | 12'-2"W x 20'-11"L
Cable Tray heights
Minimum | 8'-10"
Maximum | 9'-8"
RF Screen Height | 11'-0"
Altitude
Minimum | 98.4 feet (30 m) below sea level
Maximum | 8530 feet (2600 m) above sea level
Magnet Delivery
WxDxH in inches | 93x144x107
Weight | 25,245 lbs
Roof Hatch | 9'-0" x 9'-0"
Wall Opening | 9'-0" x 9'-0"
Delivery: min. opening, incl. Dollies
6 ft. corr = 8ft. Opening
7 ft. corr = 7 ft. Opening
8 ft. corr = 6 ft. opening
Dewar Delivery | min. 43 x 80 in door
Moving Metal Lines | Radial x Axial ft.
100-400 lbs | 3 Gauss line 13.45 x 13.45
Small Vehicles | 21 x 26
Large Vehicles | 24.5 x 30.3
Trains / elevator | See PIM
Chilled Water | Customer Supplied
Site Testing
Vibration | Optional
EMI Test | Optional
VibroAcoustic Damping Option | Required for all new sites and any upgrades where the magnet is replaced or relocated
Magnet tilt | not more than 30deg
Shipping Weights | See table 2-13 in PIM

## MR450w
MR QUICK REFERENCE - 3/3/08
Signa MR750 3.0T | Direction 5500101 rev 01
Magnet Type | Active
Magnet Mounting | Required. RF vendor supplies and installs anchors.
Table Anchored | Mobile Table
Gauss plots | refer to Gauss field tab
Cryogen Vent Size and Location from Iso | 8" Diameter 15.98" left 25.75" back
gantry cable access | Top of Magnet
Magnet weight | 24,952 lbs
conduit sizes -OC to Equip | (1) 2" & (1) 3"
RF Screen Room | 100db (10-150 MHz planewave)
Cable Kits
M7000BN | L1 = Short L2 = Short L3, L4, L5 & L6 = One length
M7000BP | L1 = Short L2 = Long L3, L4, L5 & L6 = One length
M7000BR | L1 = Long L2 = Short L3, L4, L5 & L6 = One length
voltage/phase MDP cat NO. | 480V / 3Phase
recommended ceiling height | 8'-9"
absolute min. ceiling height | 8'-2 1/2" With low ceiling height kit
absolute min. Magnet room size | 12'-2"W x 20'-11"L
Cable Tray heights
Minimum | 8'-10"
Maximum | 9'-8"
RF Screen Height | 11'-0"
Altitude
Minimum | 98.4 feet (30 m) below sea level
Maximum | 8530 feet (2600 m) above sea level
Magnet Delivery
WxDxH in inches | 88.5x94x100
Weight | 25,245 lbs
Roof Hatch | 9'-0" x 9'-0"
Wall Opening | 9'-0" x 9'-0"
Delivery: min. opening, incl. Dollies
6 ft. corr = 8ft. Opening
7 ft. corr = 7 ft. Opening
8 ft. corr = 6 ft. opening
Dewar Delivery | min. 43 x 80 in door
Moving Metal Lines | Radial x Axial ft.
100-400 lbs | 3 Gauss line 13.45 x 13.45
Small Vehicles | 21 x 26
Large Vehicles | 24.5 x 30.3
Trains / elevator | See PIM
Chilled Water | Customer Supplied
Site Testing
Vibration | Optional
EMI Test | Optional
VibroAcoustic Damping Option | Required for all new sites and any upgrades where the magnet is replaced or relocated
Magnet tilt | not more than 30deg
Shipping Weights | See table 2-13 in PIM

## 3.0T HD-HDx
MR QUICK REFERENCE - 3/3/08
3.0T Signa HDx | Direction 5159903 rev 04
3.0T Signa Excite HD | Direction 5133303 rev 05
Magnet Type | Active
Magnet Mounting | Required. RF vendor supplies and installs anchors.
Table Anchored | Mobile Table
Gauss plots | refer to Gauss plot tab
Cryogen Vent Size and Location from Iso | 8" Diameter 15.98" left 25.75" back
gantry cable access | Rear Pedistal
Magnet weight | 24,810 lbs
conduit sizes -OC to Equip | (1) 2" & (1) 3"
RF Screen Room | 100db (10-150 MHz planewave)
Cable Kits Excite HD
M3335NY | L1 = Short L2 = Short L1/L2 = Short L3, L4 & L5 = One length
M3335NZ | L1 = Short L2 = Long L1/L2 = Long L3, L4 & L5 = One length
M3335P | L1 = Long L2 = Short L1/L2 = Medium L3, L4 & L5 = One length
voltage/phase MDP cat NO. | 480V / 3Phase M3088TM
recommended ceiling height | 8'-9"
absolute min. ceiling height | 8'-7"
absolute min. Magnet room size | 11'-0"W x 20'-8"L
Magnet Delivery
WxDxH in inches | 93x144x107
Weight | 25,245 lbs
Roof Hatch | 9'-0" x 9'-0"
Wall Opening | 9'-0" x 9'-0"
Delivery: min. opening, incl. Dollies
6 ft. corr = 8ft. Opening
7 ft. corr = 7 ft. Opening
8 ft. corr = 6 ft. opening
Moving Metal Lines | Radial x Axial ft.
100-400 lbs | 3 Gauss line 13.45 x 13.45
Small Vehicles | 21 x 26
Large Vehicles | 24.5 x 30.3
Trains / elevator | See PIM
Chilled Water | Required, select GEHC supplied coldhead / gradinet chiller M3088TL = MRCC M3088TK = GWHX (heat exchanger)
Site Testing
Vibration | Optional
EMI Test | Optional
VibroAcoustic Damping Option | Required for all new Excite HDx sites and any upgrades where the magnet is replaced or relocated
Cable Kits Excite HD
M3334TD | L1 = Short L2 = Short L1/L2 = Short L3, L4 & L5 = One length
M3334TE | L1 = Short L2 = Long L1/L2 = Long L3, L4 & L5 = One length
M3334TF | L1 = Long L2 = Short L1/L2 = Medium L3, L4 & L5 = One length
Cable Kits Excite II
M3333TB | L1 = Short L2 = Short L1/L2 = Short L3 & L4 = One length L5 = Long
M3333TF | L1 = Short L2 = Long L1/L2 = Long L3 & L4 = One length L5 = Short
M3333TE | L1 = Long L2 = Short L1/L2 = Medium L3 & L4 = One length L5 = Long
Magnet tilt | not more than 30deg

## 1.5T HDe
MR QUICK REFERENCE - 3/3/08
1.5T Signa Hde | Direction 5143464 rev 04
Magnet Type | LCC Active Compact
Magnet Mounting | Required. RF vendor supplies and installs anchors.
Table Anchored | Mobile Table
Cryogen Vent Size and Location from Iso | 8" Diameter 16.00" left 13.76" back
gantry cable access | Rear Pedistal
Magnet weight | 13,116 lbs
conduit sizes -OC to Equip | (1) 2" & (1) 3"
RF Screen Room | 100db (10-100 MHz planewave)
Cable Kits Excite HD | NONE
Helium line extensions for No Equipment room cooling
inside | M20152FE
outside | M20062FY
voltage/MDC cat NO. No Equipment room or w/Equipment room & MRCC | 480V / 3Phase M3088TM
voltage/amps MDC cat NO. w/Equipment room | 480V / 3Phase 90Amp E4502SP
recommended ceiling height | 8'-9"
absolute min. ceiling height | 8'-7"
min. Magnet room size per PIM | 11'-0"W x 18'-8"L
Recommended Magnet room Area | 228.27 sqft
Magnet Delivery
WxDxH in inches | 93x144x107
Weight | 11,700 lbs
Roof Hatch | 9'-0" x 9'-0"
Wall Opening | 9'-0" x 9'-0"
Delivery: min. opening, incl. Dollies
6 ft. corr = 8ft. Opening
7 ft. corr = 7 ft. Opening
8 ft. corr = 6 ft. opening
Moving Metal Lines | Radial x Axial ft.
100-400 lbs | 3 Gauss line 7.87 x 13.45
Small Vehicles | 15.5 x 21
Large Vehicles | 18.1 x 24.5
Trains/elevator | See PIM
Chilled Water
With Equipment room cooling | Required, customer supplied Option available from GE (M3088TL for shield cooler only)
No Equipment room cooling | Required GE - M3088TL separate Cat. No.
Site Testing
Vibration | Optional
EMI Test | Optional
VibroAcoustic Damping Option | Required for all new HDe sites and any upgrades where the magnet is replaced or moved.
Dewar Sizes | LxWxH
Praxair - 250 | 32"x32"x65"
500 | 42"x42"x73"
Air Products - 250 | 32"x32"x70"
500 | 42"x42"x72"
Gauss plots | Refer to gauss plot tab

## 1.5T HD-HDx
MR QUICK REFERENCE - 3/3/08
1.5T Signa HDx | Direction 5159901 rev 04
1.5T Signa Excite HD | Direction 5133301 rev 04
Magnet Type | LCC Active Compact
Magnet Mounting | Required. RF vendor supplies and installs anchors.
Table Anchored | Mobile Table
Cryogen Vent Size and Location from Iso | 8" Diameter 16.00" left 13.76" back
gantry cable access | Rear Pedistal
Magnet weight | 13,116 lbs
conduit sizes -OC to Equip | (1) 2" & (1) 3"
RF Screen Room | 100db (10-100 MHz planewave)
Gauss plots | refer to Gauss plot tab
Cable Kits HDx
M3335NJ or M3335AA | L1 = Short L2 = Short L1/L2 = Short L3, L4 & L5 = One length
M3335NK or M3335AB | L1 = Short L2 = Long L1/L2 = Long L3, L4 & L5 = One length
M3335NL or M3335AC | L1 = Long L2 = Short L1/L2 = Medium L3, L4 & L5 = One length
voltage/amps MDC cat NO. | 480V / 3Phase 90Amp E4503AT
recommended ceiling height | 8'-9"
absolute min. ceiling height | 8'-7"
absolute min. Magnet room size | 11'-0"W x 20'-8"L
Magnet Delivery
WxDxH in inches | 88x94.5x100
Weight | 11,700 lbs
Roof Hatch | 9'-0" x 9'-0"
Wall Opening | 9'-0" x 9'-0"
Delivery: min. opening, incl. Dollies
6 ft. corr = 8ft. Opening
7 ft. corr = 7 ft. Opening
8 ft. corr = 6 ft. opening
Moving Metal Lines | Radial x Axial ft.
100-400 lbs | 3 Gauss line 7.87 x 13.45
Small Vehicles | 15.5 x 21
Large Vehicles | 18.1 x 24.5
Trains / elevator | See PIM
Chilled Water | Required, customer supplied Option available from GE (M3088TL for shield cooler only)
Site Testing
Vibration | Optional
EMI Test | Optional
VibroAcoustic Damping Option | Required for all new Excite HD or HDx sites and any upgrades where the magnet is replaced or relocated
Cable Kits Excite HD
M3134PE | L1 = Short L2 = Short L1/L2 = Short L3, L4 & L5 = One length
M3144PE | L1 = Short L2 = Long L1/L2 = Long L3, L4 & L5 = One length
M3145PE | L1 = Long L2 = Short L1/L2 = Medium L3, L4 & L5 = One length
Cable Kits Excite II
M3134PD | L1 = Short L2 = Short L1/L2 = Short L3 & L4 = One length L5 = Long
M3144PD | L1 = Short L2 = Long L1/L2 = Long L3 & L4 = One length L5 = Short
M3145PD | L1 = Long L2 = Short L1/L2 = Medium L3 & L4 = One length L5 = Long

## 0.35T Ovation
MR QUICK REFERENCE - 3/3/08
0.35T Signa Ovation 5 | Direction 5165376 rev 02
0.35T Signa Ovation | Direction 2275665 rev 09
Magnet Type | Permanent YMS
Magnet Mounting | Not required unless local code dictates. Shim plate installed unless magnetic shielding at site. GE supplied anchors
Table Anchored | Ovation 1-4 Required table rail and wearplate
For Ovation 5 there are no Table rail or ware plates.
Cryogen Vent Size and Location from Iso | NA
gantry cable access | REAR
Magnet weight | 41,888 lbs
conduit sizes -OC to Equip | (1) 2" & (1) 3"
RF Screen Room | 90db
standard cable lengths | MDP - PD = 20' IPS - SRU = 23' OC - SRU = 55' SRU - MAG = 28'
alternate cable lengths | SOI (lead time of 6mo)
voltage/amps MDC cat NO. | 480V / 3Phase 40Amp R4503K
recommended ceiling height | 8'-9"
absolute min. ceiling height | 7'-10 1/2"
absolute min. Magnet room size | 13'-2"W x 18'-9"L
Magnet Delivery
Roof Hatch | 8'-0" x 8'-0"
Wall Opening | 8'-0" x 8'-0"
Delivery: min. opening, incl. Dollies | 5ft. corr. =
6 ft. corr = 9'-0" x 9'-0"
7 ft. corr = 9'-0" x 9'-0"
8 ft. corr = 9'-0" x 9'-0"
Chilled Water | Not Required
Site Testing
Vibration | Not Required
EMI Test | Required
Moving Metal Lines | diameter ft
< 400 lbs | 9.84
Small Vehicles | 36
Large Vehicles | 42.65
Trains/elevator | See PIM

## 0.2T Profile
MR QUICK REFERENCE - 3/3/08
0.2T Signa Profile 5 | Direction 5117792 rev 02
0.2T Signa Profile | Direction 2186839 rev 04
0.2T Signa Profile HD | Direction 5196540 rev 01
Magnet Type | Permanent YMS
Magnet Mounting | Not required unless local code dictates. Shim plate installed unless magnetic shielding at site. GE supplied anchors
Table Anchored | Required
Cryogen Vent Size and Location from Iso | NA
gantry cable access | REAR
conduit sizes -OC to Equip | (1) 2" & (1) 3"
Magnet Weight | 22,046 lbs
RF Screen Room | 90db
standard cable lengths | ISP - OC = 55' ISP - SRU = 23' OC - SRU = 55' SRU - MAG = 28'
alternate cable lengths | SOI (lead time of 6mo)
voltage/amps MDC cat NO. | 480V / 3Phase 25Amp R4503PN
recommended ceiling height | 8'-9"
absolute min. ceiling height | 8'-0"
absolute min. Magnet room size | 11'-3"W x 16'-6"L
Magnet Delivery
Roof Hatch | 6'-6" x 7'-6"
Wall Opening | 5'-0" x 6'-0"
Delivery: min. opening, incl. Dollies | 5ft. corr. = 7'-2.5" x 4'-3"
6 ft. corr. = 7'-2.5" x 4'-3"
7 ft. corr. = 7'-2.5" x 4'-3"
8 ft. corr = 7'-2.5" x 4'-3"
Chilled Water | Not Required
Site Testing
Vibration | Not Required
EMI Test | Required
Moving Metal Lines | diameter ft
< 400 lbs | 9.84
Small Vehicles | 36
Large Vehicles | 42.65
Trains/elevator | See PIM

## 0.7T Openspeed
MR QUICK REFERENCE - 3/3/08
0.7T Signa Openspeed | Direction 2402692 rev 00
Magnet Type | Active
Magnet Mounting | Required. RF vendor installs anchors. GE supplies anchors
Table Anchored | Required table rail and wearplate
Cryogen Vent Size and Location from Iso | 8" Diameter 21.75" right 12.56" back
gantry cable access | REAR
Gauss plots | refer to Gauss plot tab
Magnet weight | 23,227 lbs
conduit sizes -OC to Equip | (1) 2" & (1) 3"
RF Screen Room | 100db (10-100 MHz planewave)
standard cable lengths | L1 = 14.5' L2 = 18.5' L1/L2 = 29.5' L3 = 11.5' L4 = 55' L5 = 26.5'
alternate cable lengths | SOI
voltage/amps MDC cat NO. | 480V / 3Phase 125Amp (supplied with system)
recommended ceiling height | 8'-7"
absolute min. ceiling height | 8'-7"
absolute min. Magnet room size | 13'-2"W x 19'-9"L
Magnet Delivery
Roof Hatch | 9'-0" x 9'-0"
Wall Opening | 9'-0" x 9'-0"
Delivery: min. opening, incl. Dollies
6 ft. corr = 8ft. Opening
7 ft. corr = 7 ft. Opening
8 ft. corr = 6 ft. opening
Chilled Water | Required, select GEHC supplied water chiller / M2085KC, M2085KG or M2085KE
Site Testing
Vibration | Required
EMI Test | Required
Moving Metal Lines | diameter ft
< 400 lbs | See PIM
Small Vehicles | 50
Large Vehicles | 60
Trains / elevator | See PIM

## 1.5T Twinspeed
MR QUICK REFERENCE - 3/3/08
1.5T Signa Twinspeed HDx | Direction 5159902 rev 03
1.5T Signa Twinspeed Excite HD | Direction 5133302 rev 04
Magnet Type | LCC Active
Magnet Mounting | Required. RF vendor supplies and installs anchors.
Table Anchored | Mobile Table
Cryogen Vent Size and Location from Iso | 8" Diameter 16.00" left 13.76" back
gantry cable access | Rear Pedistal
Magnet weight | 13,336 lbs
Gauss plots | refer to Gauss plot tab
conduit sizes -OC to Equip | (1) 2" & (1) 3"
RF Screen Room | 100db (10-100 MHz planewave)
Cable Kits Excite HD
M3335NP | L1 = Short L2 = Short L1/L2 = Short L3, L4 & L5 = One length
M3335NR | L1 = Short L2 = Long L1/L2 = Long L3, L4 & L5 = One length
M3335NS | L1 = Long L2 = Short L1/L2 = Medium L3, L4 & L5 = One length
voltage/amps MDC cat NO. | 480V / 3Phase 125Amp M3088TM
recommended ceiling height | 8'-9"
absolute min. ceiling height | 8'-7"
absolute min. Magnet room size | 11'-0"W x 20'-8"L
Magnet Delivery
WxDxH in inches | 93x144x107
Weight | 12,195 lbs
Roof Hatch | 9'-0" x 9'-0"
Wall Opening | 9'-0" x 9'-0"
Delivery: min. opening, incl. Dollies
6 ft. corr = 8ft. Opening
7 ft. corr = 7 ft. Opening
8 ft. corr = 6 ft. opening
Moving Metal Lines | Radial x Axial ft.
100-400 lbs | 3 Gauss line 7.87 x 13.45
Small Vehicles | 15.5 x 21
Large Vehicles | 18.1 x 24.5
Trains / elevator | See PIM
Chilled Water | Required, select GEHC supplied coldhead / gradinet chiller M3088TL = MRCC M3088TK = GWHX (heat exchanger)
Site Testing
Vibration | Optional
EMI Test | Optional
VibroAcoustic Damping Option | Required for all new Excite HD sites and any upgrades where the magnet is replaced.
Cable Kits Excite HD
M3134TE | L1 = Short L2 = Short L1/L2 = Short L3, L4 & L5 = One length
M3144TE | L1 = Short L2 = Long L1/L2 = Long L3, L4 & L5 = One length
M3145TE | L1 = Long L2 = Short L1/L2 = Medium L3, L4 & L5 = One length
Cable Kits Excite II
M3143TD | L1 = Short L2 = Short L1/L2 = Short L3 & L4 = One length L5 = Long
M3144TD | L1 = Short L2 = Long L1/L2 = Long L3 & L4 = One length L5 = Short
M3145TD | L1 = Long L2 = Short L1/L2 = Medium L3 & L4 = One length L5 = Long

## 0.5T Contour
MR QUICK REFERENCE - 3/3/08
0.5T Signa Contour | Direction 2189673 rev 04
Magnet Type | Compact Active CXK4 YMS
Magnet Mounting | Required. GE supplies kit with blocks for mounting
Table Anchored | Required
Cryogen Vent Size and Location from Iso | 6" Diameter 13.3" right 13.8" back
gantry cable access | Rear Pedistal
conduit sizes -?? to ?? -?? to ?? -?? to ??
RF Screen Room | 100db (10 - 100 MHz planewave)
standard cable lengths
alternate cable lengths | SOI
voltage/amps MDC cat NO. | 480V / 3Phase 40Amp E4503AT
recommended ceiling height
absolute min. ceiling height | 8'-9"
absolute min. Magnet room size | 12'-0"W x 20'-3"L
Magnet Delivery
Roof Hatch | 8'-0" x 8'-0"
Wall Opening | 8'-0" x 8'-0"
Delivery: min. opening, incl. Dollies | 5ft. corr. =
6 ft. corr. =
7 ft. corr. =
8 ft. corr =
Chilled Water | Not Required
Site Testing
Vibration | Not Required
EMI Test | Optional per site conditions